SPECIFICATIONS FOR INSITE DUAL DISSOLVED OXYGEN &

SUSPENDED SOLIDS MONITOR with SENSORS SUITABLE

FOR CLASS 1, DIVISION 2 AREAS

  - GENERAL REQUIREMENTS:

- The DO / SS system shall be an INSITE Model DCA dual channel Dissolved Oxygen and Suspended Solids Analyzer with one Fluorescence Dissolved Oxygen Sensor and one Suspended Solids Sensor and two barrier junction boxes.  It shall consist of the analyzer, sensors with up to 200 feet of cable, barrier junction boxes, and mounting hardware.

- The Dissolved Oxygen sensor shall be a type that does not require membranes, cartridges, fill solutions, or replaceable membrane caps as these types of designs are not acceptable. The sensor shall not be damaged by exposure to direct or indirect sunlight nor will exposure to sunlight affect the calibration.

- The Suspended Solids sensor shall be a type that measures the amount of infrared energy absorbed across a fixed sensor gap.

  - PRODUCT SPECIFICATIONS:

Analyzer

- The monitor shall be housed in a fiberglass NEMA 4X enclosure with brackets for wall or standard round handrail mounting.

- It shall have a digital display controlled by microprocessor circuitry.

- All run, programming, and calibration functions shall be accessible without having to open the enclosure.

- Provide 4 relays discrete outputs.

- Continuously display dissolved oxygen and suspended solids in mg/l.

- Provide an extended temperature, UV treated LCD digital display for continuously showing dissolved oxygen, suspended solids, and temperature readout. Programming and diagnostics are also provided through this display.

- Provide self-diagnostics for the sensor and analyzer.  Analyzer shall have error messages in the operating mode for higher or lower than normal sensor output voltage, unstable instrument circuitry, unstable output of the sensor, and weak sensor output level.

- Provide isolated RS-485 MODBUS RTU serial communication port for both sensors and relay status.

- Provide 4-20 MADC optically isolated outputs for both sensors.

- Provide standard back-lit LCD display.

- InsiteIG Model DCA analyzer

Sensors

- The Suspended Solids sensor shall be a type that measures the amount of infrared energy absorbed across a fixed sensor gap. The sensor will automatically compensate for ambient light and not be affected by changes in process color.

- The Dissolved Oxygen sensor shall be an optical type sensor that measures the fluorescence and quenching reactions of a ruthenium complex that is immobilized in a sol-gel matrix requiring no consumables and not damaged by exposure to direct or indirect sunlight.

- Ambient Temperature Range: -20ºC ≤ TAMB ≤ +50º C

- Suspended Solids Range: 0 to 30,000 mg/l

- Dissolved Oxygen Range: 0 to 25 mg/l

- Suspended Solids Resolution: 3 significant digits, 1 mg/l minimum

- Dissolved Oxygen Resolution: 0.01 mg/l below 4.0 mg/l; 0.1 mg/l above 4.0 mg/l

- Drift: Typically less than 1% per year

- Suitable for outdoor use – submerged

Class 1 Division 2 Groups A, B ,C, D, T6 per

ANSI/ISA 12.12.01-2015 and CAN/CSA C22.2 No. 213-15

- One InsiteIG Model 12 sensor and One InsiteIG Model 14 sensor

Barrier Junction Boxes

- Ambient Temperature Range: -20ºC ≤ TAMB ≤ +50º C

- Normal operating input: 8.5Vdc at 50 mA.

- Voltages in excess of 9.5Vdc may blow a non-replaceable fuse

- Currents in excess of 100 mA may blow a non-replaceable fuse

- Suitable for outdoor use. NEMA 4x (IP66)

- Two InsiteIG Model SJB-1 junction boxes