SPECIFICATIONS FOR INSITE FLUORESCENCE

DISSOLVED OXYGEN MONITOR

  - GENERAL REQUIREMENTS:

- The DO system shall be an INSITE Model 1000-10-33 single channel Dissolved Oxygen Analyzer with one fluorescence DO Sensors OR an INSITE Model 2000-1 dual channel Dissolved Oxygen Analyzer with two optical DO Sensors.  It shall consist of the monitor, sensor(s), and mounting hardware. The sensor shall be a type that does not require membranes, cartridges, fill solutions, or replaceable membrane caps as these types of designs are not acceptable. The sensor shall not be damaged by exposure to direct or indirect sunlight nor will exposure to sunlight affect the calibration.

  - PRODUCT SPECIFICATIONS:

- The monitor shall be housed in a fiberglass NEMA 4X enclosure with brackets for wall or standard round handrail mounting.

- It shall have a digital display controlled by microprocessor circuitry.

- All run, programming, and calibration functions shall be accessible without having to open the enclosure.

- Unit shall be pre-calibrated at the factory.

- Provide 4 relays discrete outputs.

- Continuously display dissolved oxygen in ppm with an accuracy of 0.02 ppm.

- Continuously display temperature in degrees C or F (user selectable).

- Provide an extended temperature, UV treated LCD digital display for continuously showing dissolved oxygen and temperature readout. Programming and diagnostics are also provided through this display.

- Provide self-diagnostics for the sensor and analyzer.  Analyzer shall have error messages in the operating mode for higher of lower than normal sensor output voltage, temperature input outside the 0-60 degree C range, and unstable instrument circuitry; error messages in the calibrate mode for unstable sensor temperature (after a 5 minute wait), unstable D.O. output of the sensor (after a 5 minute wait), and weak sensor output level.

- Provide isolated RS-485 MODBUS RTU serial communication port for dissolved oxygen, temperature, and relay status.

- Provide 4-20 MADC optically isolated outputs for both DO and Temperature.

- Provide standard back-lit LCD display.

- The sensor shall be an optical type sensor that measures the fluorescence and quenching reactions of a ruthenium complex that is immobilized in a sol-gel matrix requiring no consumables and not damaged by exposure to direct or indirect sunlight.